## Slide 1: R&S IM Spacing Sweep User Guide
Nick Lalic
VNA Software Developer
Rohde & Schwarz North America
September 2nd, 2016

## Slide 2: Typical Setup
External coupler combined input tones
Actual setup may vary with measurement needs
DUT
b2
Port2
S1
a1
Port1
S3
a3
Port3
Lower input tone
Upper input tone

## Slide 3: Settings
fcenter
flower
fupper
Start tone distance
Stop tone distance

## Slide 4: Traces
Type:
  Input
  Output
  Intermod
  Relative (normalized)
  Intercept
Feature:
  Lower
  Upper
  Major (maximum)
Order:
  3, 5, 7 or 9th order harmonics

## Slide 5: Calibration
Calibration step (optional)
Creates segmented sweep in channel with all frequencies
Application exits
Perform three power cals:
  Lower source
  Upper source
  Receiver

## Slide 6: Typical Calibration
Perform power calibration of lower input source and reference receiver at DUT input plane
Step 1
b2
Port2
S1
a1
Port1
S3
a3
Port3
Lower input tone
Upper input tone
Attach power sensor

## Slide 7: Typical Calibration
Perform power calibration of upper input source and reference receiver at DUT input plane
Step 2
b2
Port2
S1
a1
Port1
S3
a3
Port3
Lower input tone
Upper input tone
Attach power sensor

## Slide 8: Typical Calibration
Use one of the calibrated sources to perform a power calibration on the receiver
Note: There will be some imbedded loss from the thru connector used to bypass the DUT. An alternative is to calibrate the source on the unused port with the connector attached, then calibrate the receiver
Step 3
b2
Port2
S1
a1
Port1
S3
a3
Port3
Lower input tone
Upper input tone

## Slide 9: Apply traces
Restart application
Confirm settings
Skip calibration (already done)
VNA is configured for measurement

## Slide 10: Configure Display
Traces are created in new display window
Rearrange traces as necessary
Do not rename or delete: trace math depends on other traces

## Slide 11: Under the hood
The application uses the arbitrary frequency port settings to achieve measurements across tone distance
Here’s an example screenshot for the upper 3rd order intermod product with the following settings:
  Port 1 lower input tone
  Port 3 upper input tone
  Port 2 receiver (output)
  2 GHz center frequency
  1 MHz start tone spacing
  10 MHz stop tone spacing

## Slide 12: Contact
Please contact me with questions or comments
      Nick Lalic VNA Software Developer
      San Francisco, CA
      +1 (424) 200-2846 nick.lalic@rsa.rohde-schwarz.com
      http://vna.rs-us.net